| Criteria | Standard Methods Requirement | Rating |
| --- | --- | --- |
| Balance | Acceptable? |  |
| Standard Weights | Either NIST Class s or ASTM/ANSI Class 1 weights 1,2 | Yes | No |  |
| Calibration Frequency / Documentation | Calibration verification required at least once each day the balance is used.3 | Yes | No |  |
| Cleanliness, air movement, vibration | Cleanliness of balance is a must and air movement and vibration needs to be kept to a minimum1 | Yes | No |  |
| Other | Service and recalibrate annually (manufacturer representative or comparable)1 | Yes | No |  |
|  | Must be able to measure to 0.1 grams4 | Yes | No |  |
|  | Instrument manual available | Yes | No |  |
|  | Log book maintained2 | Yes | No |  |
| Comments: |
| Criteria | Standard Methods Requirement | Rating |
| Drying Oven (Suspended Solids) | Acceptable? |  |
| Temperature Recordkeeping | Temperature recorded with each use4 | Yes | No |  |
|  | Log book maintained2 | Yes | No |  |
| Calibration Frequency / Documentation | Thermometer calibrated annually with NIST traceable thermometer1,2 . Correction factor posted on thermometer / equipment1 | Yes | No |  |
| Other | Thermometer temperature in 0.5o C increments5 | Yes | No |  |
|  | Acceptable temperature range is 103o – 105o F 4 | Yes | No |  |
|  | Instrument manual available | Yes | No |  |
| Comments: |

| Criteria | Standard Methods Requirement | Rating |
| --- | --- | --- |
| pH Meter | Acceptable? |  |
| Calibration Frequency / Documentation | Calibration verification required for testing over long period of time (e.g. 12 hrs.), or after a large number of samples (every 10 samples)3 | Yes | No |  |
|  | Logbook maintained2 | Yes | No |  |
| Minimum of 2 point calibration | Calibration per manufacturer specification and calibration buffers must bracket anticipated result7 | Yes | No |  |
| Slope Documentation / Acceptability | Slope acceptable range indicated on benchsheet2 | Yes | No |  |
| Buffer Expiration Date | Buffers must not be expired | Yes | No |  |
| Other | Instrument manual available | Yes | No |  |
|  | Teflon covered magnetic stirrer or equivalent for mixing8 | Yes | No |  |
| Comments: |
| Criteria | Standard Methods Requirement | Rating |
| Dissolved Oxygen Meter | Acceptable? |  |
| Calibration Method | Air or known DO calibration method10 | Yes | No |  |
|  | Calibration per manufacturer specification10 | Yes | No |  |
| Calibration Frequency / Documentation | Logbook maintained2 | Yes | No |  |
|  | Calibration verification required at least once each day the meter is used.3 | Yes | No |  |
| Other | Small to no bubble present under membrane (must be smaller than the lead in number 2 pencil)11 | Yes | No |  |
|  | Instrument manual available | Yes | No |  |
| Comments: |

| Criteria | Standard Methods Requirement | Rating |
| --- | --- | --- |
| Incubator (CBOD/ E-Coli) | Acceptable? |  |
| Temperature Recordkeeping | Temperature checked / recorded twice daily for each shelf in use1 | Yes | No |  |
|  | Acceptable temperature range (CBOD) is 20o C ±1.0 o 12 | Yes | No |  |
|  | Acceptable temperature range (E-Coli) is 35o C ±0.5 o 22 | Yes | No |  |
|  | Logbook maintained 2 | Yes | No |  |
| Temperature Calibration / Documentation | Thermometer calibrated annually with NIST traceable thermometer1, 2 | Yes | No |  |
|  | Temperature correction information posted on incubator1 | Yes | No |  |
| E-Coli can use multiple tubes (five 20 ml or ten 10 ml), or mfg’s multi-well tray | E-coli Ultraviolet lamp (365 nm wave length, 6 W bulb) 23 | Yes | No |  |
| Other | Instrument manual available | Yes | No |  |
|  | Temperature Log (thermometer reads to 0.5 Celsius).1 | Yes | No |  |
| Comments: |
| Criteria | Standard Methods Requirement | Rating |
| Refrigerator | Acceptable? |  |
| Temperature Recordkeeping | Temperature Log (thermometer reads to 0.5 Celsius).5 | Yes | No |  |
| Temperature Calibration / Documentation | Thermometer calibrated annually with NIST traceable thermometer 1, 2 | Yes | No |  |
| Other | Thermometer held in water bath. 1 | Yes | No |  |
|  | Refrigerator temperature ≤6o Celsius.13 | Yes | No |  |
|  | Do not store volatile solvents, food, or beverages.14 | Yes | No |  |
| Comments: |

| Criteria | Standard Methods Requirement | Rating |
| --- | --- | --- |
| Chlorine Meter | Acceptable? |  |
| Calibration Frequency / Documentation | pH / millivolt meter read to 0.1 mV15 | Yes | No |  |
|  | Calibration verification required for testing over long period of time (e.g. 12 hrs.), or after a large number of samples (every 10 samples)3 | Yes | No |  |
| Calibration Method | Calibration using three iodate solutions 0.2, 1.0, 5.0 milliliters or calibration per manufacturer specification16 | Yes | No |  |
|  | Standards used for calibration not expired | Yes | No |  |
| Slope Documentation / Acceptability | Calibration curve (acceptable slope) | Yes | No |  |
| Other | Electrode free of deposits and foreign material | Yes | No |  |
|  | Log book being maintained.2 | Yes | No |  |
|  | Instrument manual available | Yes | No |  |
| Comments: |
| Criteria | Standard Methods Requirement | Rating |
| Ammonia Meter | Acceptable? |  |
| Calibration Frequency / Documentation | Calibration verification required for testing over long period of time (e.g. 12 hrs.), or after a large number of samples (every 10 samples)3 | Yes | No |  |
|  | Log book being maintained2 | Yes | No |  |
| Slope acceptability | Verify calibration slope is acceptable (per mfg. spec.). | Yes | No |  |
| Calibration Method | Standards used for calibration (3 ammonia solutions of 10 mg/l, 1 mg/l, and 0.1 mg/l) or per mfg. spec.17 | Yes | No |  |
|  | Standards used for calibration not expired | Yes | No |  |
| Other | Electrode free of deposits and foreign material | Yes | No |  |
|  | Teflon covered magnetic stirrer or equivalent for mixing18 | Yes | No |  |
|  | Instrument manual available | Yes | No |  |
| Comments: |

| Criteria | Standard Methods Requirement | Rating |
| --- | --- | --- |
| Sample Collection/Handling | Acceptable? |  |
| Sample Labeling | Samples container labeled (description, date, time, preservative added, initialed).19 | Yes | No |  |
| Chain of Custody | Chain of custody (description, date, time, signature).19 | Yes | No |  |
| Other | Composite samples refrigerated during sample collection14 | Yes | No |  |
|  | Equipment blanks utilized14 | Yes | No |  |
|  | SOP for cleaning of sampling equipment | Yes | No |  |
|  | Logbook being maintained2 | Yes | No |  |
| Comments: |
| Criteria | Standard Methods Requirement | Rating |
| Desiccator | Acceptable? |  |
| General criteria | Properly working seals. | Yes | No |  |
|  | Desiccant fresh (blue color) | Yes | No |  |
| Documentation | Log book being maintained2 | Yes | No |  |
| Comments: |
| Criteria | Standard Methods Requirement | Rating |
| Bench sheets | Acceptable? |  |
| General criteria | Date(s)2 | Yes | No |  |
|  | Analyst initials2 | Yes | No |  |
|  | Blue or black ink pen2 | Yes | No |  |
|  | Calibration information2 | Yes | No |  |
|  | Equations, calculations, units for all measurements, notations, and results present2 | Yes | No |  |
|  | Corrections, single line through, initialed and dated2 | Yes | No |  |
| Comments: |

| Criteria | Standard Methods Requirement | Rating |
| --- | --- | --- |
| Hot Water Bath (Fecal Coliform/E. Coli) | Acceptable? |  |
| Temperature Recordkeeping | Temperature Log (thermometer reads 0.2o C)21 | Yes | No |  |
|  | Incubator temperature 44.5o C ± 0.2o 21/24 |  |  |  |
| Temperature Calibration / Documentation | Thermometer calibrated annually with NIST traceable thermometer 1, 2 | Yes | No |  |
|  | Log book being maintained 2 | Yes | No |  |
| Water Level | Thermometer total immersion or partial (line on thermometer to ID immersion depth) 1, 5 | Yes | No |  |
| Comments: |
| Criteria | Standard Methods Requirement | Rating |
| Autoclaves/Steam Sterilizers | Acceptable? |  |
| All apparatus utilized is adequately sterilized before use | Sterilizing temperature 121o C 25 | Yes | No |  |
|  | 10 to 30 minutes time based on material being sterilized26 | Yes | No |  |
| Documentation | Verify the autoclave temperature weekly by using a maximum registering thermometer (MRT) to confirm that 121°C has been reached as measured in the exhaust. 1 | Yes | No |  |
|  | Date, contents, sterilization time and temperature, total time in autoclave, and analyst’s initials should be recorded each time the autoclave is used 1 | Yes | No |  |
| Temperature Calibration / Documentation | Thermometer calibrated annually with NIST traceable thermometer 1,2 | Yes | No |  |
|  | Log book being maintained 2 | Yes | No |  |
| Performance Checks | Test monthly for efficacy using a biological such as commercially available Geobacillus stearothermophilus in spore strips, suspensions, or capsules 1 | Yes | No |  |
| Comments: |

| Criteria | Standard Methods Requirement | Rating |
| --- | --- | --- |
| Final Effluent Temperature Monitoring | Acceptable? |  |
| General Criteria | Thermometer calibrated annually with NIST traceable thermometer 1,2 | Yes | No |  |
|  | Thermometer reads in increments of at least 0.1o C 5 | Yes | No |  |
|  | Log book being maintained 2 | Yes | No |  |
| Comments: |
| Number of Criteria Rated: | Acceptable |  |
|  | Marginal |  |
|  | Unacceptable |  |
|  | Total Number of Areas Rated |  |
| Acceptable Ratings – No action required (recommend SOP’s written or updated, perform DMRQA’s for all onsite analysis, recommend voluntary lab analyst certification, written response not required). |
| Marginal Ratings – Improvements required, written response required (recommend SOP’s be written or updated, recommend they perform DMRQA’s for all onsite analysis, recommend voluntary lab analyst certification, require deficiencies to be addressed in written response). |
| Unsatisfactory Rating - Improvements required, written response required, NOV issued (recommend SOP’s be written or updated, recommend they perform DMRQA’s for all onsite analysis, recommend voluntary lab analyst certification, require deficiencies to be addressed in written response to NOV). |
| Consider recommending PAI Audit from DES when: | >60% of ratings are Marginal >45% of ratings are a combination of Marginal or Unacceptable >30% of ratings are Unacceptable |

| Notation of Referenced Method |
| --- |
| 1 | Method 9020-B, Item 3 | 14 | Method 1060A, Item 1 |
| 2 | Method 1020-A, Item 1 | 15 | Method 4500-Cl I, Item 2 |
| 3 | Method 1020-B, Item 10 | 16 | Method 4500-Cl I, Item 4 |
| 4 | Method 2540-B, Item 2 | 17 | Method 4500-NH3 D, Item 4 |
| 5 | Method 2550-B, Item 1 | 18 | Method 4500-NH3 D, Item 2 |
| 6 | Method 1020-A, Item 1 | 19 | Method 1060-B, Item 2 |
| 7 | Method 4500-H B, Item 4 | 20 | Method 1060-B, Item 1 |
| 8 | Method 4500-H B, Item 2 | 21 | Method 9222D, Item 1 |
| 9 | Method 1020-B, Item 2 | 22 | Method 9223 B, Item 2 |
| 10 | Method 4500-O B, Item 3 | 23 | Method 9223 B, Item 3 |
| 11 | Method 4500-O G, Item 3 | 24 | Method 1603, Item 2 |
| 12 | Method 5210-B, Item 5 | 25 | Method 9030-B, Item 3 |
| 13 | CFR 136.3, Table II | 26 | Method 9020 B, Table IV |

Equipment Logbook Content - all maintenance performed on a piece of equipment should be documented in the logbook.  This should include parts replacement and routine maintenance activities.  Entries should include date, maintenance performed and initials of person making entry.

| Preservation and Holding Times |
| --- |
| Parameter | Container | Min. Sample Size (mL) | Sample Type | Preservation | Maximum Storage Time |
|  |  |  |  |  | Recommended | Regulatory |
| BOD / CBOD | P, G | 1000 | G, C | Refrigerate ≤6o C | 6h | 48h |
| TSS | P, G | 200 | G, C | Refrigerate ≤6o C | 7 d | 7 d |
| pH | P, G | 50 | G | Analyze immediately | 0.25h | 0.25 h |
| NH3-N | P, G | 500 | G, C | Analyze as soon as possible or add H2SO4 to pH <2, Refrigerate ≤6o C | 7 d | 28 d |
| TRC | P, G | 500 | G | Analyze immediately | 0.25h | 0.25 h |
| DO (electrode) | G, BOD Bottle | 300 | G | Analyze immediately | 0.25h | 0.25 h |
| Temperature | P, G | -- | G | Analyze immediately | 0.25h | 0.25 h |
| Metals, general | P, G | 1000 | G, C | For dissolved filter immediately and add HNO3 to pH <2 | 6 months | 6 months |
| Purgeables by purge and trap | G (PTFE lined lid) | 40 (X2) | G | HCl to pH<2, Refrigerate ≤6o C | 7 d | 14 d |
| Base/Neutrals and acids | G (solvent rinsed or baked) | 1000 | C, G | Refrigerate ≤6o C | 7 d | 7 days until extraction 40 days after extraction |
| Pesticides | G (PTFE lined lid) | 1000 | C | Refrigerate ≤6o C | 7 d | 7 days until extraction 40 days after extraction |
| Fecal Coliform / E-Coli | G, P (Sterilized) | 100 | G | Refrigerate ≤10o C If chlorine present, add sodium thiosulfate tablet | 6 hrs transport Start analysis within 2 hrs of receipt in lab. |
| Oil and Grease | G | 1000 | G | HCl or H2SO4 to pH <2, Refrigerate ≤6o C | 28 d | 28 d |

| Approved Standard Methods |
| --- |
| CBOD / BOD 5 Day | Std Methods 5210-B |
| Ammonia, Selective Electrode Method | Std Methods 4500-NH3 D |
| Total Residual Chlorine, DPD Colorimetric Method | Std Methods 4500-Cl G |
| Total Suspended Solids, Dried at 103-105 °C | Std Methods 2540-D |
| Dissolved Oxygen, Membrane Electrode Method | Std Method 4500-O G |
| pH, Electrometric Method | Std Methods 4500-H+ B |
| Fecal Coliform, Membrane Filter Procedure | Std Methods 9222D |
| Escherichia Coli, Enzyme Substrate Test | Std Method 9223B |
| Escherichia Coli Membrane Filtration Procedure | EPA Method 1603 |
| Oil and Grease | USEPA 1664A or Std Methods 5520B |
| Metals, general | USEPA 200, Std Methods 3111B or C, or 3120B |
| Volatiles (Purgeables by purge and trap) | USEPA 6210, Std Methods 624 |
| Semi-Volatiles (Base/Neutrals and acids) | USEPA 6410, Std Methods 625 |
| Pesticides | USEPA 6410 and 6630, Std Methods 608 |